## RiskWorkSheet
Risk Assessment Form
Group Name: | Application of Cyclic Stress for Treatment of Subglottic Stenosis |  | Members: |  |  | Severity | Exposure | Probablity
|  |  | Taylor Hughes |  |  | Catastrophic | Frequent | Probable
|  |  | Brian Dallesasse |  |  | Serious | Occasional | Possible
|  |  | Kyle Sachdev |  |  | Slight | Remote | Unlikely
[6 empty] | Minimal | None | Negligible
Item | User | Task | Hazard Category | Hazard | Cause/Failure Mode | Severity | Exposure | Probability | Risk Level | Method of Risk Reduction/Mitigation | Severity | Exposure | Probability | Risk Level | Person Responsible | Date | Comments
1.1 | Patient | Operate on Patient/Sterilize Equipment | Biological/Health | Infection | Device not thoroughly kept to sterilization proceeding insertion | Serious | Remote | Unlikely | High | Proper sterilization before insertion, reliable clinician team | Minimal | None | Negligible | Medium | Kyle Sachdev | 3.8.17 | Verification w/ Dr. Paniello that hospital staff is reliable
1.2 | Patient | Operate Device | Mechanical | Inability to pass air to vocal cords for speech | Balloon deflates and fills air passage gap | Minimal | Occasional | Possible | Medium | Addition of solid structure for balloon to inflate/deflate around | Minimal | None | Negligible | Low | Taylor Hughes | 2.28.17 | Loss of airflow to vocal cords will not directly impair health of patient, but rather indirectly by blocking speech ability which could effect their communicate other problems that may arise with the treatment
1.3 | Patient | Operate Device | Mechanical | Tubing rupture, power board exposed to liquid | Tubing material not strong to handle sustained liquid pressure | Slight | Remote | Unlikely | Medium | Stronger tubing as obtained from client | Minimal | None | Negligible | Low | Brian Dallesasse | 3.23.17 | Circuity exposure to liquid is dangerous and should not be in close proximity to people
1.4 | Patient | Operate Device | Mechanical | Hemoptysis - Tracheal abrasion/bleeding | Balloon slips and applies pressure to non-target area | Serious | Remote | Possible | High | Back support anchor to prevent slippage | Slight | None | Unlikely | Medium | Taylor Hughes | 3.20.17 | Anchor piece is reflective of cricoid cartilage for better support
1.5 | Patient | Operate Device | Mechanical | Delay of Medical Care | Actuator/syringe interface detached | Slight | Remote | Possible | Low | End of actuator fits into top part of syringe plunger, add super glue to secure | Minimal | Remote | Unlikely | Low | Kyle Sachdev | 4.11.17 | Patient condition does not improve/poor health sustained if mechanism to support this treatment fails
1.6 | Patient | Operate Device | Mechanical | Respiration impairment | Device slips and falls into tracheostomy tube air path | Catastrophic | Remote | Unlikely | High | Back support anchor to prevent slippage | Slight | Remote | Unlikely | High | Taylor Hughes | 3.20.17 | Anchor piece is reflective of cricoid cartilage for better support. The device to fall all the way to the trache is unlikely, but a risk nonetheless. The support should prevent this unlikely occurance
1.7 | Patient | Operate on Patient/Sterilize Equipment | Eronomic/Human Factors | Cut to neck | Surgeon not carefull w/ knife when making incision point for tubing | Slight | Remote | Negligible | High | Properly trained physician | Minimal | None | Negligible | Medium | Brian Dallesasse | 3.2.17 | Our client, Dr. Paniello, is very much a trustworthy clinician and is able to handle procedures of this type
1.8 | Patient | Operate Device | Mechanical | Balloon rupture, water/saline flow into lungs | Applied pressure greater than max pressure witheld by balloon material | Catastrophic | Remote | Unlikely | High | Arduino code set to not allow linear actuator to push syringe beyond set point; the syringe/actuator interface is enclosed and cannot be accidentally altered from outside forces | Slight | None | Negligible | Medium | Kyle Sachdev | 4.11.17 | Arduino code is preset and cannot be altered once device is in use, so applied pressure will remain constant throughout treatment
2.1 | Doctor | Operate on Patient | Biological/Health | Unsanitary conditions | Careless with sterilization procedures | Slight | Remote | Negligible | Medium | Wear gloves & mask | Minimal | None | Negligible | Low | Kyle Sachdev | 3.2.17 | Our client practices standard sanitary procedures
3.1 | All Users | Work on Equipment | Laser | Laser used for laser incision can damage eyes | Careless use by laser operator | Serious | Remote | Negligible | High | Wear goggles | Slight | Remote | Negligible | Low | Brian Dallesasse | 2.20.17 | Brian witnessed safety precaution taken among hospital staff in room during real procedure

## Lists
Severity | Exposure | Probability | Risk | Users | Hazard Category | Hazard
Catastrophic | Frequent | Probable | High | All Users | Biological/Health
Serious | Occasional | Possible | Medium | Doctor | Mechanical
Slight | Remote | Unlikely | Low | Nurse | Electrical
Minimal | None | Negligible |  | Technician | Slip/Trip/Fall
|  |  |  | Fabricator | Eronomic/Human Factors
|  |  |  | Housekeeping | Fire/Explosion
|  |  |  | Paramedic | Heat/Temperature
|  |  |  | Nurse Aide | Noise/Vibration
|  |  |  | Service Technician | Ingress/Egress
|  |  |  | Patient | Material Handling
|  |  |  | Family | Environmental/Industrial Hygene
|  |  |  | Public | Ventilation/Confined Space
Chemicals/Gases
Fluid Pressure
Radiation
Laser
Wastes